## Subjects
subject_id | session_date | age | sex | height_cm | weight_kg | bmi | hypertension | medication | medication_name | smoking | finger | cuff_arm | sbp_baseline_1 | dbp_baseline_1 | sbp_baseline_2 | dbp_baseline_2 | sbp_baseline_mean | dbp_baseline_mean | sbp_post | dbp_post | csv_file | duration_s | fs_hz | signal_quality_notes | general_notes | arm_circumference_cm | inter_arm_diff_sbp | hours_since_caffeine | hours_since_exercise | hr_baseline_1 | hr_baseline_2 | hr_baseline_3 | bp_category | cv_fold | protocol_version | quality_tier | training_eligible | quality_flag_reason | pi_ir_pct | medication_dose_mg | medication_taken_today_hr_ago | years_since_diagnosis | bp_category_diagnosed | classification_note
S001 | 2026-04-29 | 23 | M | 182 | 68 |  | n | n | n | n | middle | left | 105 | 72 | 108 | 69 |  |  | 110 | 78 | Danh01_20260429_051348.csv | 373.8 | 104 |  | Đo lúc 5.am, ngồi yên 5 phút trước đo | [10 empty] | A | Y |  | 2.45 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S002 | 2026-04-29 | 23 | M | 162 | 60 |  | n | n | n | y | middle | left | 97 | 70 | 99 | 67 |  |  | 106 | 61 | Quoc_01_20260429_104215.csv | 579.2 | 104 |  | Đo lúc mới ngủ dậy, ngôì yên 5 phút trước đo | [10 empty] | A | Y |  | 1.53 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S003 | 2026-04-28 | 23 | M | 182 | 68 |  | n | n | n | n | middle | left | [6 empty] | 99 | 67 | Self_test 01_20260428_030211.csv | 362.2 | 104 | [12 empty] | A | Y |  | 1.85 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S004 | 2026-04-29 | 19 | F | 152 | 45 |  | n | n | n | n | middle | left | 93 | 67 | 99 | 75 |  |  | 104 | 80 | Minh_01_20260429_145144.csv | 462.4 | 91 |  | Chưa ăn sáng; ngồi im | [10 empty] | C | N | PI low 0.58% (sat reject 0.5%); F/19; calibration anchor only | 0.58 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S005 | 2026-04-29 | 19 | M | 167 | 62 |  | n | n | n | n | middle | left | 124 | 81 | 118 | 77 |  |  | 112 | 82 | Sub_3_20260429_150828.csv | 705.2 | 104 |  | Ăn sáng; ngồi im khi đo | [10 empty] | B | Y | fs drift 80.7Hz vs claim 104; fixed sau khi sua train_ppg_bp.py | 1.38 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S006 | 2026-04-29 | 23 | M | 174 | 62 |  | n | n | n | n | middle | right | 124 | 85 | 130 | 82 |  |  | 123 | 83 | CH_1_20260429_162631.csv | 691.7 | 104 |  | Ngồi im khi do; ăn uống đầy đủ | [10 empty] | A | Y |  | 1.47 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S007 | 2026-05-01 | 23 | M | 170 | 63 |  | n | n | n | n | middle | left | 102 | 67 | 103 | 64 |  |  | 108 | 67 | DUy_01_20260501_103156.csv | 374 | 105 | [12 empty] | A | Y |  | 1.56 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S008 | 2026-05-01 | 32 | M | 161 | 63 |  | n | n | n | y | index | left | 127 | 75 | 127 | 77 |  |  | 101 | 79 | Hung_01_20260501_175617.csv | 274.1 | 104 | Excellent — AC% IR = 24.92%, DC = 83,561 (in target) | Subject 32 tuổi, smoker. BP_post chênh -26 mmHg vs baseline (127→101) → có thể white-coat effect lần đầu. Cần re-measure session 2 với protocol v1.1 (3 readings, drop lần 1). | [7 empty] | Elevated |  | v1.0 | A | Y |  | 24.92 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S009 | 2026-05-01 | 32 | M | 173 | 78 |  | n | n | n | n | index | left | 111 | 67 | 111 | 78 |  |  | 115 | 72 | Huy_01_20260501_182144.csv | 292.3 | 105 | Excellent — AC% IR = 32.02% (kỷ lục dataset), DC = 83,149 | Subject 32t, BMI 26.1 Overweight. BP stable 111/73→115/72 (chênh +4 OK). Cùng tuổi Hung_01 nhưng khác BMI/smoking/BP — pair stratify analysis ideal. | [7 empty] | Normal |  | v1.0 | A | Y |  | 32.02 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S010 | 2026-05-01 | 19 | M | 170 | 71 |  | n | n | n | n | index | left | 124 | 72 | 128 | 75 |  |  | 118 | 75 | Kin_01_20260501_184341.csv | 291.3 | 105 | Excellent — AC% IR = 5.31%, DC = 79,581 (borderline, dưới target 80k 0.5%) | Subject 19t M, BMI 24.6 Normal, no smoking. BP stable 124/72 → 118/75 (chênh -6 OK). Elevated category (SBP 120-129). | [7 empty] | Elevated |  | v1.0 | A | Y |  | 5.31 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S011 | 2026-05-01 | 33 | F | 153 | 43 |  | n | n | n | n | middle | left | 96 | 70 | 96 | 69 |  |  | 90 | 67 | tram_01_20260501_190404_cleaned.csv | 279 | 111 | Original file có 6 ESP32 disconnects (12-29s gaps, 92.5s dead time). Cleaned via clean_gap_csv.py: 31000 samples, 279s clean @ 111Hz. AC% raw = 14.05%. | Subject 33t Female (NỮ THỨ 2 dataset!), BMI 18.4 Underweight. 3 attempts do ESP32 disconnect: cTram_01 (159s partial), Tram_01 (67s partial), tram_01 (full 60 batches, dùng cleaned). BP stable 96/70 → 90/67. | [7 empty] | Normal |  | v1.0 | B | Y | 6 ESP32 disconnects, dead time 92.5s/389s = 24%. Cleaned version dùng được. | 14.05 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S012 | 2026-05-01 | 45 | F | 150 | 62 |  | n | n | n | n | index | left | 115 | 80 |  |  |  |  | 127 | 88 | coBa_01_20260501_195111.csv | 298.1 | 104 | Excellent — AC% IR 14.03%, DC 80316 (in target). 0 disconnects, no cleaning needed. | Subject 45t F (FIRST ≥40t!), BMI 27.6 Overweight. Stage 1 HTN measured (DBP 80-88). BP chênh baseline→post: +12 SBP/+8 DBP (borderline, < 15 threshold). Mean(115,127)=121 / Mean(80,88)=84 dùng làm label. Quan trọng nhất dataset: fill 3 gaps cùng lúc (HTN + Female + ≥40t). | [7 empty] | Stage 1 HTN |  | v1.0 | A | Y |  | 14.03 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S013 | 2026-05-05 | 32 | M | 170 | 75 | 26 | n | n | n | n | index | right | 126 | 96 | 128 | 94 | 127 | 95 | 126 | 93 | Viet_01_20260505_194040.csv | 306.1 | 100 | 5min full session. WiFi/upload stall mid-session: 7 gaps total dead time 72s (22.4%). clean_gap_csv concat 8 segments + boundary trim 700 samples → 30614 samples @ 100Hz, 306.1s effective. Cleaned: Viet_01_20260505_194040_cleaned.csv. fs_real intra-segment 100Hz exact. Pred SBP 106.00 / DBP 70.80 (errors -21.0/-24.2 vs cuff). Pred raw≈cleaned (signal robust). | Stage 2 HTN subject — FIRST in dataset (was 0%). Contralateral protocol CONFIRMED (cuff right + index finger LEFT) per user 2026-05-05. Note: opposite direction to user default (cuff left + trỏ right) — cả 2 đều contralateral hợp lệ. Metadata finger originally typo "middle" → corrected to "index" in CSV header + here. WiFi/upload stall is firmware uploadQueue (10-batch buffer) limitation, not data quality issue. ML errors -21/-24 expected for out-of-distribution Stage 2 HTN trên model degenerate (Moulaeifard 2025 baseline ~18-25 mmHg). | [7 empty] | Stage 2 HTN |  | v1.1 | B | Y | WiFi/upload stall 7 gaps mid-session (22.4% dead time, recovered via clean_gap_csv concat). 306s effective > 120s protocol min. fs 100Hz exact post-bug-fix. | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S014 | 2026-05-05 | 37 | F | 157 | 52 | 21.1 | n | n | n | n | index | left | 108 | 70 | 101 | 64 | 104.5 | 67 | 100 | 73 | _archived_protocol_violation/Phuong_01_20260505_195608.csv | 304 | 100 | 5min full session. 1 disconnect 11.9s at batch 49 (3.9% dead time) + LED auto-recalibration after disconnect (RED DC +8.6%, ratio R/IR 0.967→1.031). clean_gap_csv concat 2 segments → 30400 samples @ 100Hz, 304.0s. POST-only pred SBP 126.20 vs PRE-only 117.00 (Δ +9.2 mmHg shift due to LED change). Cleaned ≈ PRE pred (PRE 82% sample weight). Pred SBP 116.80 / DBP 73.90 (errors +12.8/+6.9 vs cuff). | PROTOCOL VIOLATION same-arm (cuff TRÁI + ngón trỏ TRÁI) — confirmed by user 2026-05-05. Reactive hyperemia post-cuff may confound PPG amplitude/morphology in first 1-2 min. EXCLUDED from training (training_eligible=N). Files ARCHIVED to collected_data/_archived_protocol_violation/. Action: re-recruit Phuong session 1 with contralateral protocol (cuff right + finger left). Female ≥30 demographic value (rare, only 2/14 unique training subjects: coBa 45F + Phuong 37F). LED recalib mid-session quantified +9.2 mmHg shift PRE vs POST — bằng chứng features không hoàn toàn LED-invariant. | [7 empty] | Normal |  | v1.1 | C | N | Same-arm protocol violation (cuff left + index finger left). Methodological flaw not recoverable by signal cleaning. Excluded from training. Re-recruit needed contralateral. Files archived 2026-05-05. | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S015 | 2026-05-03 | 19 | F | 152 | 45 |  | n | n | n | n | index | left | 87 | 60 | 94 | 57 |  |  | 93 | 62 | Minh_02_20260503_153026.csv | 304.7 | 100 | Session 2 (3/5). 2 network hiccups (~11s each, total 22.5s) cleaned. Post-clean: 304.7s @ 100Hz, AC% IR 32.09% (great improvement vs S1 13.49%). DC stable 86458 throughout — finger never lifted. Cleaned file: Minh_02_*_cleaned.csv | Session 2 of Minh (S005). Cách session 1 (29/4) 4 ngày, cùng khung giờ ~15:00 ✓ protocol v1.1. BP drop S1→S2: -5 SBP, -12 DBP (96/71→91/59) ⇒ borderline hypotensive. Subject technique improved (AC% 2.4× better). Network 2 hiccups handled by firmware queue. FIRST session 2 in dataset → critical for testing per-subject calibration cross-session. | [7 empty] | Normal |  | v1.1 | B | Y | 2 network hiccups ~11s, cleaned to 304.7s continuous. 8% gap (vs 25% Sub_3 accepted). | 32.09 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S016 | 2026-05-03 | 19 | M | 167 | 62 |  | n | n | n | n | index | left | 111 | 77 | 110 | 73 |  |  | 118 | 75 | Sub_3_02_20260503_155516.csv | 292.6 | 100 | Session 2 (3/5) — CLEAN signal: 0 disconnects, AC% IR 6.98%, DC 80865 (in target). fs_median=100Hz exact (sensor đúng nominal, không drift 111Hz như session 1). KHÔNG cần clean_gap_csv. | Session 2 of Sub_3 (S006). Cách session 1 (29/4 15:08) 4 ngày, khung giờ 15:55 ✓ protocol v1.1. BP DROP S1→S2 (Stage 1 HTN→Normal): SBP 124→111 (-13), DBP 81→75 (-6). Có thể stress/anxiety effect lần đầu. Mean(baseline,post): SBP 114.5, DBP 75 → Elevated. Signal session 2 cleaner hẳn (S1 có 8 disconnects 100s dead, S2 không gap nào). | [7 empty] | Normal |  | v1.1 | A | Y |  | 6.98 |  |  |  |  | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S017 | 2026-05-05 | 23 | M | 182 | 68 | 20.5 | n | n | n | n | index | left | 105 | 67 |  |  | 105 | 67 | 109 | 76 | Danh_02_20260505_190801.csv | 292.6 | 100 | 5min full session. fs_real=100.00Hz EXACT (drift 0.0% — first file post timestamp truncation bug fix log_ppg_local.py:211). NO disconnects (>1s gaps). clean_gap_csv: no cleaning needed. IR mean 75780, RED mean 79772. HR pred 85.5 BPM. Pred SBP 106.50 / DBP 70.30 (errors +1.50/+3.30 vs cuff baseline 105/67) — AAMI compliant on single point. FIRST file with 100Hz exact timestamps — all prior 13 self-collected files had fs_real=111Hz drift due to old bug. | Session 2 of Danh (S001). Cách session 1 (29/4 5am) 6 ngày, khung giờ shifted 5am→7pm (DIFFERENT time of day, protocol v1.1 prefers same time ±1h — limitation). BP very stable cross-session: S1 106/70, S2 105/67 (Δ -1/-3 mmHg, within Omron ±3 noise). Cross-session pred test: pred S1=107.5/71.5, pred S2=106.5/70.3 (Δpred -1.0/-1.2) → CONSISTENT direction with cuff Δ; per-subject calibration offset will be reliable. Contralateral protocol CONFIRMED (cuff left + index finger RIGHT) — user default 2026-05-05. Third session 2 in dataset (after S005-2 Minh, S006-2 Sub_3). | [7 empty] | Normal |  | v1.1 | A | Y | [6 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S018 | 2026-05-05 | 37 | F | 157 | 52 | 21.1 | n | n | n | n | index | left | 100 | 66 | 102 | 68 |  | 67 | 97 | 74 | Phuong_01v2_20260505_214139.csv | 288.8 | 100 | CLEANEST file in dataset to-date. fs_real=100.00Hz EXACT, 0 disconnects, only 3 small gaps (>100ms) at batch boundaries. NO LED auto-recalibration mid-session (smooth Q1→Q4 R/IR drift 0.969→0.960 = -1% over 5min, no abrupt jump). IR mean 87712, RED mean 84552 (higher DC than archived S030: IR 79850, RED 78131). HR pred 82.4 BPM. Pred SBP 109.50 / DBP 70.30 (errors +8.50/+3.30 vs cuff baseline 101/67). Cleaning: not needed. | RE-RECORDING of Phuong_01 session 1 (S030 ARCHIVED for same-arm protocol violation). Contralateral protocol CONFIRMED per user 2026-05-05: cuff TRÁI + index finger TAY PHẢI (user default). Metadata finger initially typo "middle" → corrected to "index" in CSV header + here. THESIS-RELEVANT: Same subject same evening (S030 19:56 vs S015 21:41, ~1h45m apart) — cuff BP nearly identical (S030 104/67→101/67 baseline, Δ -3 SBP within Omron noise). PRED SBP shifted from 116.80 (same-arm S030) to 109.50 (contralateral S015) = Δ -7.30 mmHg. Quantifies REACTIVE HYPEREMIA BIAS contribution to PPG-based BP estimation when same-arm protocol used. Strong evidence for thesis Limitations citing AHA 2017 / ARMS Trial 2024 contralateral requirement. Female ≥30 demographic value preserved. | [7 empty] | Normal |  | v1.1 | A | Y | [6 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S019 | 2026-05-06 | 51 | M | 168 | 70 | 24.8 | y | y | n | n | index | left |  |  |  |  | 134 | 93 | 136 | 90 | Triet_01_20260506_182107_stripped_cleaned.csv | [11 empty] | Stage 2 HTN |  | v1.1 | B | Y | low_perfusion_index_PI_0.34pct | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S020 | 2026-05-06 | 48 | F | 162 | 66 | 25.1 | n | n | n | n | index | left |  |  |  |  | 135 | 89 | 142 | 92 | Trang_20260506_183827_stripped_cleaned.csv | [11 empty] | Stage 1 HTN |  | v1.1 | B | Y | low_PI_0.36pct;ac_healthy;regression_to_mean_err_-23.7_SBP | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S021 | 2026-05-07 | 23 | M | 162 | 60 | 22.9 | n | n | n | y | index | left |  |  |  |  | 100 | 63 | 103 | 57 | Quoc_02_20260507_152350_cleaned.csv | [11 empty] | Normal |  | v1.1 | A | Y | clean_post_K_fix_fs100Hz_exact_drift_session_2 | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S022 | 2026-05-08 | 23 | M | 182 | 68 | 20.5 | n | n | n | n | index | left |  |  |  |  | 103 | 71 | 113 | 79 | Danh_04_20260508_063934_cleaned.csv | 305 | 100 | [9 empty] | Normal |  | v1.1 | A | Y | 2_batch_led_recalib_stripped_drift_session_4_fs100Hz_exact | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S023 | 2026-05-08 | 23 | M | 174 | 62 | 20.5 | n | n | n | n | index | left |  |  |  |  | 112 | 78 | 111 | 78 | CH_20260508_134606_cleaned.csv | 283 | 100 | [9 empty] | Normal |  | v1.1 | A | Y | 5_batch_led_recalib_stripped_drift_session_2_BP_DROP_127_to_112_white_coat_S1 | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S024 | 2026-05-08 | 23 | M | 170 | 64 | 22.1 | n | n | n | n | index | left | 103 | 70 | 103 | 76 | 103 | 73 | 99 | 64 | DUy_20260508_191456_cleaned.csv | 260.9 | 100 | [9 empty] | Normal |  | v1.1 | A | Y | 5_batch_led_recalib_stripped_drift_session_2_S3_used_S2_archived_white_coat_pre113_post105_S3_pre103_matches_S1_anchor | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S025 | 2026-05-08 | 32 | M | 173 | 78 | 26.1 | n | n | n | n | index | left | 126 | 74 | 131 | 80 | 129 | 77 | 129 | 75 | Huy_20260508_205239_cleaned.csv | 296 | 100 | [9 empty] | Elevated |  | v1.1 | A | Y | led_recalib_same_intensity_4_disconnects_cleaned_22pct_dead_time_genuine_BP_rise_S1_111_to_S2_129_pre_eq_post_NOT_white_coat | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S026 | 2026-05-08 | 37 | F | 157 | 52 | 21.1 | n | n | n | n | index | left | 102 | 64 | 92 | 61 |  |  | 97 | 63 | Phuong_20260508_211411_cleaned.csv | 269.5 | 100 | [9 empty] | Normal |  | v1.1 | A | Y | 5_batch_led_recalib_stripped_drift_session_2_stable_BP_pre_eq_post_short_gap_2.98d_SpO2_R_below_0.4_threshold_14pct_valid | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S027 | 2026-05-10 | 23 | M | 174 | 62 | 20.5 | n | n | n | n | index | left |  |  |  |  | 123 | 75 | 128 | 75 | CH_20260510_151202_cleaned.csv | 271.5 | 100 | [9 empty] | Elevated |  | v1.1 | A | Y | 5_batch_led_recalib_stripped_drift_session_3_BP_variability_S1_127_S2_112_S3_123_NOT_white_coat_averaged_anchor_breakthrough_SBP_drift_0.58_mmHg | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation).
S028 | 2026-05-11 | 51 | M | 168 | 70 | 24.8 | y | y | n | n | index | left |  |  |  |  | 130 | 86 | 131 | 93 | Triet_02_20260511_191356_stripped_cleaned.csv | 268.7 | 100 | Tier 1 recruit S2; no disconnects; fs=100 exact post-strip | Sustained Stage 1 reading S1(134/93 Stage2)->S2(130/86 Stage1), DBP drop -7 mmHg | [7 empty] | Stage 1 reading |  | v1.1 | B | Y | Low PI expected (S1 had 0.34%) | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation across days).
S029 | 2026-05-11 | 48 | F | 162 | 66 | 25.1 | n | n | n | n | index | left |  |  |  |  | 132 | 88 | 137 | 96 | Trang_20260511_193148_stripped_cleaned.csv | 271 | 100 | Tier 1 recruit S2; 8 disconnects 28.6s stripped; 9 segments | Stage 1 reading sustained S1(135/89)->S2(132/88); pre->post within-session white-coat +5/+8 SBP/DBP | [7 empty] | Stage 1 reading |  | v1.1 | B | Y | Low PI expected (S1 had 0.36%); 10.6% dead time from disconnects | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation across days).
S042 | 2026-05-11 | 32 | M | 170 | 75 | 26 | n | n |  | n | index | left |  |  |  |  | 130 | 85 | 122 | 88 | Viet_20260511_214922_stripped_cleaned.csv | 268.7 | 100 | Tier 1 recruit S2; no disconnects; fs=100 exact post-strip | Stage 2 (DBP=95)->Stage 1 reading (DBP=85), DBP drop -10 mmHg same pattern as Triet. cuff_arm switched right(S1)->left(S2) still contralateral | [7 empty] | Stage 1 reading |  | v1.1 | B | Y | cuff_arm switched between sessions; inter-arm BP difference possible confound +/-5-10 mmHg | [5 empty] | Single-session BP reading per AHA 2017. NOT HTN diagnosis (requires multi-session sustained elevation across days).
[13 empty] | 126 | 81 | 133 | 88
Note on AA1: Vòng tay PHẢI (Omron fit 22-32 cm)
Note on AB1: Chênh BP 2 tay (chỉ Session 1; >10 note, >20 exclude)
Note on AC1: Giờ kể từ lần uống cà phê/trà gần nhất
Note on AD1: Giờ kể từ lần vận động mạnh gần nhất
Note on AE1: HR Omron lần 1 (drop)
Note on AF1: HR Omron lần 2
Note on AG1: HR Omron lần 3
Note on AH1: IMPORTANT (per supervisor 2026-05-10): This column shows BP READING CATEGORY at single session per AHA 2017 / ESH 2018 classification. It is NOT a clinical HTN diagnosis. HTN diagnosis requires: - 2+ sessions on 2+ different days - Sustained elevation across visits - Or 24h ABPM, or physician confirmation Single elevated reading = elevated reading category, NOT a diagnosis. Example: CH S1=127 (Stage 1 reading) but S2=112 (Normal), S3=123 (Elevated) - BP variability, NOT HTN.
Note on AI1: LOSO fold assignment (1-15) cho training
Note on AJ1: v1.0 hoặc v1.1 (track schema migration)
Note on AK1: A=clean train+calib, B=salvageable, C=calib only
Note on AL1: Y/N - dung de train ML
Note on AM1: Ly do flag
Note on AN1: PI tu eval (AC IR / DC IR * 100)
Note on AO1: Liều thuốc HA (vd 5, 10/12.5 cho combo)
Note on AP1: So gio ke tu lan uong thuoc gan nhat
Note on AQ1: So nam ke tu chan doan HTN
Note on AR1: None / Stage 1 HTN / Stage 2 HTN (theo diagnosis cu, KHONG dung lam ML label)

## BP_Reference
AHA Blood Pressure Categories (2017+)
Category | SBP (mmHg) | DBP (mmHg) | Target % thesis recruit
Normotensive | <120 | và <80 | ~30%
Elevated | 120-129 | và <80 | ~10%
Hypertension Stage 1 | 130-139 | hoặc 80-89 | ~25%
Hypertension Stage 2 | ≥140 | hoặc ≥90 | ~30%
Hypertensive crisis | >180 | và/hoặc >120 | ~5% (nếu có)
Tại sao cần diversity?
Dataset PPG-BP Liang 2018 có ~80% normotensive, 20% hypertensive.
→ Self-collected nên ưu tiên Stage 1+2 (gia đình lớn tuổi, người dùng thuốc HA).
→ Hypertensive samples sẽ giúp model generalize tốt hơn cho high-risk users.

## Protocol_Checklist
AHA Standard Protocol cho đo BP + PPG
TRƯỚC ĐO (5 phút)
☐ Subject ngồi yên trong phòng yên tĩnh
☐ KHÔNG nói chuyện, không sử dụng điện thoại
☐ KHÔNG vận động mạnh < 30 phút
☐ KHÔNG uống cà phê/trà < 30 phút
☐ KHÔNG hút thuốc < 30 phút
☐ Tay đặt ngang tim, lòng bàn tay ngửa lên
☐ Chân đặt phẳng trên sàn, không bắt chéo
ĐO BP cuff Omron HEM-7143T1 (2 lần × 30s)
☐ Cuff đặt 1-2cm trên khuỷu tay (tay không thuận)
☐ Cuff vừa khít, lùa được 2 ngón tay vào
☐ Press START Omron → đo lần 1 → ghi sbp_baseline_1, dbp_baseline_1
☐ Đợi 1 phút giữa 2 lần đo
☐ Đo lần 2 → ghi sbp_baseline_2, dbp_baseline_2
☐ Excel auto-tính sbp_baseline_mean, dbp_baseline_mean
ĐO PPG (5 phút streaming)
☐ Run laptop server: python log_ppg_local.py
☐ Nhập subject_id + metadata + sbp/dbp baseline khi prompt
☐ MAX30102 đặt ngón giữa hoặc ngón trỏ tay khác
☐ Tay không cử động, đặt yên trên bàn
☐ Press ENTER trên console khi xong (~5 phút)
ĐO BP POST (1 lần, sau khi xong PPG)
☐ Đo Omron lần 3 → ghi sbp_post, dbp_post
☐ Mục đích: check BP drift trong session
LƯU + BACKUP
☐ CSV PPG đã save automatic vào collected_data/
☐ Update Excel này: csv_file name, duration_s, fs_hz, notes
☐ Backup folder collected_data/ + Excel lên Google Drive cuối ngày

## Schema_v1.1_Notes
SCHEMA v1.1 MIGRATION NOTES (01/05/2026)
LƯU Ý QUAN TRỌNG:
1. Hàng 2-22 (S001-S014?) ĐO THEO PROTOCOL v1.0 (5 phút resting, có thể chỉ 2 lần BP).
→ Giữ nguyên data, KHÔNG retroactively fix. Note ở cột protocol_version = "v1.0".
2. Hàng MỚI từ row 23+ phải theo PROTOCOL v1.1:
- Resting 10 phút (không phải 5)
- Đo BP 3 lần, drop lần 1, MEAN(lần 2, lần 3)
- 2 sessions cách nhau 2-7 ngày, cùng khung giờ ±1h
- Inter-arm BP screening Session 1
- protocol_version = "v1.1"
3. Formula mới cho sbp_baseline_mean (rows v1.1):
= ROUND( (sbp_baseline_2 + sbp_baseline_3) / 2, 0 ) -- skip lần 1
4. AHA BP categories (cột bp_category):
- Normal: SBP < 120 AND DBP < 80
- Elevated: SBP 120-129 AND DBP < 80
- Stage 1: SBP 130-139 OR DBP 80-89
- Stage 2: SBP >= 140 OR DBP >= 90
5. cv_fold: assign sau khi xong recruit (1-15 LOSO)
Reference: QUY_TRINH_THU_MAU.md v1.1 (01/05/2026)